## Introduction

## The Ohaus USB Interface Kit is a unique solution to the problem of connecting a balance to a computer using a Universal Serial Bus (USB). USB devices are categorized into classes such as disk drives, digital cameras, printers, etc. Balances do not have a commonly used class so the Ohaus USB interface uses a generic interface based on the RS232 serial standard.

## Data sent from the balance to a computer is in USB format. The USB data is directed to a virtual port. This port then appears as an RS232 port to the application program.

## When sending a command from a computer to the balance, the application program sends a command to the virtual port as if it were an RS232 port. The computer then directs the command from the virtual port to the computers USB connector where the balance is connected. The Adapter Kit receives the USB signal and reacts to the command.

The USB Interface kit includes a CD with the software drivers to create the required virtual port on the computer.

## System Requirements

- PC running Windows 98®, Windows 98SE®, Windows ME®, Windows 2000® or Windows XP®

- Available USB port (Type A, 4-pin, female)

USB Output (Continued)

2

Data sent from the Scout Pro is in standard ASCII format terminated with a CRLF (carriage return-line feed). The Output string is as follows:

[weight]	10 characters (right justified)

[space]	1 character

[unit]	5 characters max (left justified)

[space]	1 character

[stability ind.] 	1 character:  ? when unstable, blank when stable

[space]	1 character

[legend]	10 character:  TOTAL, hh:mm:ss (interval), etc.

[CR]	1 character

[LF]	1 character

Output examples: (Note: * and _ indicate spaces)

*****192.21_g	- Manual, Continuous or On-Stable printing

*******0.01_g_?	- Unstable reading

*******0.01_g___00:00:00	- Interval Printing (15 sec interval)

*****176.30_g_?_00:00:15	- Unstable reading at 15 seconds

*****192.08_g___00:00:30	- Stable reading at 30 seconds

*****192.21_g

*****207.80_g_TOTAL	- Totalize Mode (manual printing only)

7

Balance Setup (Continued)

Setup Definitions

#### PRINT menu:

Stable - On	Only stable data is output

Auto Print - Continuous	Data is output continuously

Auto Print - On Stable	Data is automatically output upon stability

Auto Print - (xx) sec	Data is output every (xx) seconds

Auto Print - Off	Data is only output when PRINT button on balance is pressed

USB menu:

USB - Off	Turns interface off

Baud, Parity, Handshake	Set parameters to match printer or computer

## Operation

USB Output

With the Adapter installed the balance will operate in one of the following three ways.

- USB = on, Auto Print = off, Stable = on or off
Pressing PRINT will send the display data. If Stable is set on, the balance will wait for a stable reading before sending data.

- USB = on, Auto Print = on, Stable = on or off
The balance will automatically send data based on the settings in the Menu. If Stable is set on only stable will be sent.

- RS232 = off
The interface is turned off. During battery operation, turning the USB Interface off saves significant power.

Interface Installation

6

1.  	2.

3.  	4.

## USB Connection

The USB cable terminates with a 4-pin, male, USB Type A connector.

- Power on the computer and verify that the USB port is enabled and working properly.

- Plug the cable’s USB connector into the computer’s USB connector. Windows® should detect a USB device and the New Hardware Wizard will be initialized.

3

## Virtual Port Software Installation

## Insert the supplied CD into the CD drive.

- Different versions of Windows® have slightly different steps to load the driver that is on the CD. In all versions the New Hardware Wizard guides you through the required steps to select the driver that is located on the CD.

Example of Windows 2000 Hardware Wizard

- After clicking Finish, the virtual port should be ready for use.

Windows® typically adds the virtual port in sequence after the highest number COM port. For example, on PC’s equipped with up to 4 COM ports, the virtual port will be COM5.

When using the USB interface with programs that limit the number of COM port designations (e.g. Ohaus MassTracker allows only COM1, 2, 3, & 4), it may be necessary to assign one of these port numbers to the new virtual port. This can be done in the Port Settings of the Device Manager utility, found in the Windows Control Panel.

Balance Setup

4

The Scout Pro interface is preset to communicate using the following settings:

2400 baud, 7 bit, no parity, no handshake

If different settings are desired, it will be necessary to change either the Scout Pro settings, or the computer settings.

Upon installation of the interface, Scout Pro will recognize the USB Adapter and add PRINT and USB menus to the structure. Configure the Scout Pro to operate as desired. Refer to the Scout Pro User Manual for assistance in using the menus.

CAL	SETUP	UNITS   	PRINT	USB	END

Stable		USB

-	On	-	On

-	Off	-	Off

A-Print		Baud

-	Cont	-	600

-	On.stbl	-	1200

-	5sec	-	2400

-	15sec	-	4800

-	30sec	-	9600

-	60sec	-	19200

-	Off

End		Parity

-	Yes	-	7-even

-	No	-	7-odd

-	7-none

-	8-none

Handshake

-	None

-	Xon-Xoff

-	RTS-CTS

END

-	Yes

-	No

5

## Operation (Continued)

USB Input

Scout Pro will respond to various commands sent via the interface adapter. The balance will respond to the following commands when terminated with a CR or CRLF.

?	prints current mode

0A	turns Auto-print off

SA	Auto-print on, prints on stability

CA	Continuous Auto-print

(n)A	Auto-print on 1 to 3600 second intervals (n = 1 to 3600)

C	initiates Span calibration

L	initiates Linearity calibration

0M	gram mode

1M	ounce mode

2M	troy ounce mode

3M	pennyweight mode

4M	parts counting mode

5M	pound mode

T	same as pressing ON-ZERO

V	print software version

(Esc)R	resets balance to factory defaults

P	same function as pressing the PRINT key

LE	prints last error code (e.g. Err 0)

0S	print unstable data

1S	print stable data only

© Ohaus Corporation 2003, all rights reserved.

P/N 80250955 rev.-  *80250975*			Printed in China

SCOUT Pro USB Interface Kit

INSTRUCTION MANUAL

This Interface Kit is for use only with the Ohaus Scout Pro. This accessory has been tested and complies with the approvals listed in the Scout Pro Instruction Manual.

Ohaus Corporation

19A Chapin Road

P.O. Box 2033

Pine Brook, NJ  USA

Tel: (973) 377-9000

www.ohaus.com